| Clone | TNF1172 |
| --- | --- |
| Source | Mouse Monoclonal |
| Cat # | PM214-6ml RTU PM214-3ml RTU CM214-0.5ml Conc CM214-0.5ml Conc HAM214-6ml RTU HAM214-3ml RTU |
| Regulatory Status | IVD |

TNF-alpha (TNF1172)

Intended Use:

This antibody is intended for use to qualitatively identify TNF alpha (Tumor Necrosis Factor-α) antigen by light microscopy in formalin fixed, paraffin embedded tissue sections using immunohistochemical detection methodology. Interpretation of any positive or negative staining must be complemented with the evaluation of proper controls and must be made within the context of the patient’s clinical history and other diagnostic tests. A qualified pathologist must perform evaluation of the test.

Summary and Explanation:

This MAb recognizes human 17-26kDa protein, which is identified as cytokine TNF-alpha (Tumor Necrosis Factor-alpha). TNF-alpha can be expressed as a 17kDa free molecule, or as a 26kDa membrane protein. TNF-alpha is a protein secreted by lipopolysaccharide-stimulated macrophages, and causes tumor necrosis when injected into tumor bearing mice. TNF alpha causes cytolysis of certain transformed cells, being synergistic with interferon gamma in its cytotoxicity. Although it has little effect on many cultured normal human cells, TNF alpha appears to be directly toxic to vascular endothelial cells. Other actions of TNF alpha include stimulating growth of human fibroblasts and other cell lines, activating polymorphonuclear neutrophils and osteoclasts, and induction of interleukin 1, prostaglandin E2 and collagenase production. TNF alpha is currently being evaluated in treatment of certain cancers and AIDS Related Complex.

Isotype: 	Mouse IgM/kappa

Reagent Provided: 
  	Concentrated format: Antibody to TNF α is diluted in antibody diluent, with 1% bovine serum 
 				albumin (BSA) and 0.05% sodium azide (NaN3). Recommended dilutions: 
 				1:50- 1: 100.The antibody dilution and protocol may vary depending on 
 				the specimen preparation and specific application. Optimal conditions should be 
 				determined by individual laboratory.

Pre-diluted format:  	PathnSitu ready to use antibodies are pre tittered to optimal staining 
  		           		conditions. Further dilution may loose the activity and may yield to sub 
 				optimal staining.

Storage Recommendations: Store at 2°-8°C. Do not use after expiration date provided on the vial.

Staining Recommendations: 
 	Antigen Retrieval Solution: Use TRIS-EDTA Buffer (PathnSitu Cat # PS009) as antigen retrieval 
 					solution Heat Retrieval Method: Retrieve sections under steam pressure 
 					for 15 min using PathnSitu’s MERS (Multi Epitope Retrieval System) then 
 					allow solution to cool for 10 minutes then transfer tissue sections/slides to 
 					distilled water.

Primary Antibody: 		Cover the tissue sections with primary antibody and incubate for 30 
 				min at room temperature when used PathnSitu PolyExcel Detection 
 				System.

Detection System: 	Refer to PathnSitu PolyExcel detection system protocol or manufacturer’s detection kit staining protocol when used other vendor detection system.

Cellular Localization: 	Cytoplasm and extracellular (Secreted)

Positive Control: 		Histiocytoma

Troubleshooting: 	Follow the antibody specific protocol recommendations according to data sheet provided. If unusual results occur, contact PathnSitu Technical Support at 040-2701 5544 or techsupport@pathnsitu.com.

Limitations and Warranty: 	There are no warranties, expressed or implied, which extend beyond this 
                                      	description. PathnSitu is not liable for property damage, personal injury, or 
                                   		economic loss caused by this product.

Bibliography: 		1. Bebok Z; Markus B; Nemeth P. Prognostic relevance of transforming growth 
 				    factor alpha (TGF- alpha) and tumor necrosis factor alpha (TNF-alpha) 
 				    detected in breast cancer tissues by immunohistochemistry. Breast Cancer 
 				    Research and Treatment, 1994, 29(3):229-35.